Important! Safety Notice and Recall of Certain Fluke Digital Clamp Meters

Fluke is voluntarily recalling certain 333, 334, 335, 336 and 337 Digital Clamp Meters that were manufactured between 29 January 2009 and 27 February, 2009. The models and serial numbers affected are listed in the table below. If you own one of these clamp meters, please stop using it and send it back to Fluke for replacement. If you are not the primary user of the Fluke 33X Series Clamp Meter, please pass this notice along to the appropriate people within your organization

| Model Family | Starting Serial No. | Ending Serial No. |
| --- | --- | --- |
| Fluke-337 (All models which start with 337, including model 337A and 337CWG) | 96070000 | 98890607 |
| Fluke-336 (All models which start with 336, including model 336A and 336 CWG) | 96220000 | 98900400 |
| Fluke-335 (All models which start with 335, including model 335A and 336CWG) | 97010000 | 98860104 |
| Fluke-334 (All models which start with 334, including model 335A and 336CWG) | 97010000 | 98870101 |
| Fluke-333 (All models which start with 333, including model 333A and 336CWG) | 97010000 | 98880240 |

PLEASE NOTE:  Models with the suffix “R” or the prefix “S” are NOT included in this recall.

Description of the problem

Through quality assurance tests, we have discovered a potential short circuit connection on the circuit board. This may lead to inaccurate voltage readings including a low or no voltage reading on a circuit energized with a hazardous voltage. This finding could create a hazardous situation if the user were to contact live voltage based on an erroneous instrument reading.

Actions

1. Stop using your affected 33X series clamp meter(s) even if you have not experienced problems.

2. Verify your clamp meter is affected using the serial numbers provided above.

3. Return your clamp meter for a free replacement. You can return your clamp meter by:

- Returning your meter to your place of purchase.  They will Supply a replacement meter free of charge from their stock holding, or

OR

- Completing the attached form and fax to Associated Enterprises at 977-1-261744.  We will advise you how to return the affected meter. Once the stock is available, we will supply a replacement meter to you.

OR

- Calling Mr. Sailesh at 977-1- 4250934 between 9:00 am and 5 pm Monday to Friday for further details

4. Please do not send test leads, cases or other accessories.

We are undertaking this voluntary recall because we place tremendous importance on our customer’s safety and the reputation of our products.  Please accept our apologies for any inconvenience caused by this action.

Fluke Digital Camp Meter Return Form

Model number:	___________________________________________________________

Serial number:	___________________________________________________________

Please complete this form and return it with your clamp meter.

Date:____________________ Name:__________________________________________

Company Name: __________________________________________________________

Shipping Address (No PO Boxes Please): _______________________________________

_________________________________________________________________________

Country:______________________________ Post Code:___________________________

Phone Number: ____________________________________________________________

Email Address: _____________________________________________________________

Send to:

Attention: 33X Returns (Mr Sailesh)

Associated Enterprises

Kamal Arcade 3 FL

Putalisadak Kathmandu

Nepal