Fluke 1735 Firmware Update Release Notes

Revision 01.08
Fixes minor calibration bug introduced by V01.07

Revision 01.07
Added support for 6000A iFlex set introduced for 174x series 

Revision 01.06
Add missing minima and maxima to all values in “Power” logging mode. Add plausibility check on timestamp read from RTC.

Revision 01.05
Fixes wrong logging results for THD. Firmware used different algorithms for THD in logging and online modes. 

Revision 01.04
Update to accommodate new compiler. 

Revision 01.03
Fixes “peep-hole” in the flash memory detection algorithm: 

Revision 01.01
Supports new AMD flash memory. New devices come in 48pin housing and require a change in micro programs for chip erase, sector erase, and byte program. Revision 01.02

Revision 01.00
This is the base revision for this document.